| SOP for: (check one) Process Hazardous Chemical Hazard Class of Chemicals |  |  |
| --- | --- | --- |
| Name of Procedure: Using Liquid Nitrogen |  |  |
| Applies to Lab (name): Chamberland Research Group | SCI Room #: 313 |  |
| Prepared By: Stephen Chamberland | Date: 8/31/12 | Revision: 8/31/12 |

# Special Notifications:

Liquid nitrogen is commonly used in the Chamberland lab to cool the insulated trap for the high vacuum pump to prevent vapors from reaching the vacuum pump oil.  It is used rarely as a refrigerant in tandem with an organic solvent to prepare a low-temperature cold bath or to execute a freeze-pump-thaw cycle to degas a solvent or solution.

# Hazard Description:

Liquid nitrogen is extremely cold (-195 °C). Contact with liquids at this temperature will cause serious tissue burn. When spilled on a surface, liquid nitrogen tends to spread as far as the quantity of liquid spilled and the physical confines of the area permit. It can cool large areas. The vapors are also extremely cold and can produce burns. Exposure to these cold gases, which is too brief to affect the skin of the face or hands, may affect delicate tissues, such as the eyes. Stand clear of boiling and splashing liquid and the cold vapors that are released. Boiling and splashing always occur when filling a warm container with liquid nitrogen or when inserting warm objects into the cold liquid. Always perform these operations slowly to minimize the splashing and boiling.

Do not touch un-insulated surfaces cooled by liquid nitrogen or by the vapor of liquid nitrogen. The flesh could become stuck to the cold surfaces and will rip as you remove it from the cold surfaces. Splashing may occur when liquid nitrogen is dispensed into a warmer container. Extreme care should be taken to protect sensitive tissues such as eyes. Never allow any unprotected part of the body to touch uninsulated pipes or vessels, which contain cryogenic fluids. The extremely cold metal will cause the flesh to stick fast and tear when one attempts to withdraw from it. Even nonmetallic materials are dangerous to touch at these low temperatures. Use tongs to immerse and remove objects from cryogenic liquids. In addition to the hazards of frostbite or flesh sticking to cold materials, objects that are soft and pliable at room temperature, such as rubber or plastics, are easily broken because they turn brittle at low temperatures and may break when stressed.

Although nitrogen is nontoxic and inert, it can act as a simple asphyxiant by displacing the oxygen in air to levels below that required to support life. Inhalation of nitrogen in excessive amounts can cause dizziness, nausea, vomiting, loss of consciousness, and death. Death may result from errors in judgment, confusion, or loss of consciousness that prevents self-rescue. At low oxygen concentrations, unconsciousness and death may occur in seconds and without warning. Personnel, including rescue workers, should not enter areas where the oxygen concentration is below 19.5%, unless provided with a self-contained breathing apparatus or air-line respirator.

# Protective Equipment:

- Protective wrap-around goggles

- Lab coat

- Long pants

- Close-toed shoes

- Loose fitting thermal insulated or leather gloves (NOTE:  Gloves should be loose-fitting so they are able to be quickly removed if cryogenic liquid is spilled on them. Insulated gloves are not made to permit the hands to be put into a cryogenic liquid. They will only provide short-term protection from accidental contact with the liquid.

# Engineering Controls:

Store and use with adequate ventilation. Do not store in a confined space. Cryogenic containers are equipped with pressure relief devices to control internal pressure. Under normal conditions these containers will periodically vent product. Do not plug, remove, or tamper with pressure relief device.  Cryogenic containers should be filled no more than 80% of capacity to protect against thermal expansion.

# Special Handling or Storage Requirements:

Store and use with adequate ventilation. Do not store in a confined space. Cryogenic containers are equipped with pressure relief devices to control internal pressure. Under normal conditions these containers will periodically vent product. Do not plug, remove, or tamper with pressure relief device.  Cryogenic containers should be filled no more than 80% of capacity to protect against thermal expansion.

# Spill Response Procedure:

A small amount of spilled liquid nitrogen will rapidly boil away, so there is no need for cleanup.  A large amount of spilled liquid nitrogen presents a cold and asphyxiant hazard, so avoid the puddle of liquid nitrogen and move to fresh air.

# Decontamination/Waste Disposal:

N/A

# Special Emergency Procedures:

In the event of leaks to the liquid nitrogen tanks: Evacuate all personnel from the affected area. Notify

Jeff Wilcox and/or Emil Babik. Do NOT attempt repairs.

INHALATION: Persons suffering from lack of oxygen should be moved to fresh air. Call 911.

SKIN CONTACT: Remove any clothing that may restrict circulation to frozen area. Do not rub frozen parts as tissue

damage may result. As soon as practical, place the affected area in a warm water bath which has a temperature not to

exceed 105°F (40°C). Never use dry heat. Call 911.

Frozen tissue is painless and appears waxy with a possible yellow color. It will become swollen, painful, and prone to

infection when thawed. If the frozen part of the body has been thawed, cover the area with dry sterile dressing with a

large bulky protective covering, pending medical care. In case of massive exposure, remove clothing while showering

with warm water. Call 911.

EYE CONTACT: For exposure to liquid, immediately warm frostbite area with warm water (not to

exceed 105°F). Call 911.

Procedure may be performed after normal working hours (M-F, 8-5)?	Yes    No

Procedure may be left unattended?		Yes     No

# Step by Step Procedure to get Liquid Nitrogen from its storage container:

- Take the 1L dewar with cap to the NMR room and sign the LN2 log book indicating how much LN2 that will be taken from the large dewar.  Also bring the orange insulated gloves.

- Set the 1L dewar on a wooden stool and insert the latex rubber hose emanating from the large dewar into the 1L dewar so that it is almost to the bottom.

- Very slowly, open the lever on the liquid nitrogen dewar to introduce liquid nitrogen into the small dewar.  Nitrogen gas vapors will come at first, but liquid will begin to flow as the tubing cools.  If the lever is opened too much, liquid nitrogen is likely to splash out of the small dewar and/or the rubber tubing will flail about.

- Fill the small dewar with liquid until it is at least ½ to ¾ full. Cryogenic containers should be filled no more than 80% of capacity to protect against thermal expansion.

- Using the insulated glove, turn off the flow of liquid nitrogen from the large storage tank.  The valve may be too cold to touch without the glove.

- Hold the 1L dewar by its metal handle (handle may be cold) and remove the stool from beneath the dewar.  Lower the 1L dewar so that the now frozen and brittle rubber tubing is no longer in the dewar.

- Secure the cap to the top of the 1L dewar.

- If the low oxygen alarm goes off (it is likely to do so), fan the device until the alarm shuts off.  If you unplug the device while fanning around the area to dissipate the nitrogen gas, plug it back in afterward.  When the oxygen sensor is plugged back in, the alarm always goes off initially and then stops.  If the alarm does not stop, then the sensor is still registering too little oxygen in the vicinity.

Approved By (PI) : Stephen Chamberland, Ph.D.					Date: 8/31/12

PI Signature: ___________________________________________________		Date: ______________

EH&S Signature: ________________________________________________		Date: ______________