| SOP for: (check one) Process Hazardous Chemical Hazard Class of Chemicals |  |  |
| --- | --- | --- |
| Name of Procedure: Using the Heat Gun |  |  |
| Applies to Lab (name): Chamberland Research Group | SCI Room #: 313 |  |
| Prepared By: Stephen Chamberland | Date: 9/5/12 | Revision: 9/5/12 |

# Special Notifications:

The heat gun is used in the laboratory to transfer thermal energy from one item to another, usually to save time.

# Hazard Description:

The heating element of most heat guns consists of exposed wire which typically becomes red hot during use. In addition, the on-off switch on a heating gun is not spark-free. Therefore, heat guns must never be used near flammable liquids and vapors. The power cords on heating guns often fray near the handle. Heat guns must be continually inspected for this problem and repaired immediately if frayed.  Heat guns are a source of extremely high temperature, and the aluminum shield near the nozzle can burn the skin if touched.

# Protective Equipment:

- Lab coat

- Protective wrap-around goggles

- Thermally resistant gloves

- Long pants

- Close-toed shoes

# Engineering Controls:

Remove flammable solvents from the area where the heat gun will be used.  Do not use the heat gun for any purpose other than those specified in this SOP.  Do not operate the heat gun on a setting higher than 5.  Temperatures above 200 °C are reached above the setting of 5, which are not necessary for any purpose listed herein.

Temperatures produced at different settings by the Chamberland Research Group heat gun were determined on 9/5/12.  To measure these temperatures, an alcohol thermometer was held horizontally with the thermometer bulb centered approximately 1 inch above the nozzle.  After the dial was adjusted to each setting, the temperature was recorded after approximately two minutes or until the temperature stabilized.

Heat gun setting	Temperature (°C)

1	20

1.5	25

1.75	41

2	55

2.5	71

3	109

3.5	130

4	144

4.5	170

5	194

# Special Handling or Storage Requirements:

Do not touch the metal nozzle when hot.  Do not operate the heat gun near flammable solvents.  Do not allow the cord to dangle on the floor.  Hold the heat gun by its handle and never by the cord.  Do not use the heat gun as a heat source for distillation of a flammable solvent.

# Spill Response Procedure:

N/A

# Decontamination/Waste Disposal:

N/A

# Special Emergency Procedures:

If any exposed skin touches the hot nozzle of the heat gun or if prolonged contact with a hot air stream causes a burn, run warm water over the affected area for fifteen minutes before administering first aid.  Consult a physician for severe burns.

Procedure may be performed after normal working hours (M-F, 8-5)?	Yes    No

Procedure may be left unattended?		Yes     No

# Step by Step Procedure:

The heat gun may be used to heat TLC plates after they have been dipped in a developing stain, to warm the ends of plastic tubing to assist connection to glass tips, such as on the double manifold, and to heat the bump trap (only when necessary) on the rotary evaporator when removing a solvent with a boiling point of over 100 °C.  This last use is common when removing water, DMSO, and DMF on the rotovap.

To visualize glass-backed TLC plates when using a developing stain (do not use the heat gun to visualize plastic or aluminum-backed plates)

- Dip the TLC plate in the necessary stain, and then drain the excess on the lip of the jar and then on a paper towel.

- Turn on the heat gun and hold the TLC plate with forceps in the stream of hot air until it can be determined that spots are visible.

- Turn off the heat gun and let the TLC plate cool before handling.  The tips of the forceps will be hot.

To warm the ends of plastic tubing to assist the joining of plastic tubing to glass

- Turn on the heat gun and hold the end of the plastic tubing in the stream of hot air until the tubing becomes more pliable.

- Turn off the heat gun.

- Fit the tubing (use a hot glove) on the glass tip.

To “chase” a solvent condensed in the bump trap when removing a high-boiling solvent on the rotovap.

- Turn on the heat gun and hold the nozzle about 1 inch away from the bump trap.  Move the nozzle around a bit to avoid forming a hot spot.

- When the rotary evaporation is judged to be complete, turn off the heat gun, allow the bump trap to cool, and pour any excess solvent into the appropriate waste container.

Approved By (PI) :  Stephen Chamberland, Ph.D.					Date: 9/5/12

PI Signature: ___________________________________________________		Date: ______________

EH&S Signature: ________________________________________________		Date: ______________