| SOP for: (check one) Process Hazardous Chemical Hazard Class of Chemicals |  |  |
| --- | --- | --- |
| Name of Procedure: Using the Rotary Evaporator |  |  |
| Applies to Lab (name): Chamberland Research Group | SCI Room #: 313 |  |
| Prepared By: James Siegenthaler, Stephen Chamberland | Date: 11/30/11 | Revision: 8/29/12 |

# Special Notifications:

The machine is fragile.  Before using: turn on the circulating chiller so that the cooling solution is at 4 °C, ensure that the cold finger has isopropyl alcohol and dry ice (if dry ice is needed, add it slowly), and fill the heat bath with deionized water until it is about 1-2" from the top.

# Hazard Description:

Dry ice can cause burns to the skin. A solution of dry ice and isopropyl alcohol reaches -78 °C and can cause burns to the skin. Dry ice should be handled carefully to avoid cold burns. Various organic solvents will be evaporated and re-condensed.  Based on what is being used, hazards must be addressed based on the chemical's SDS.

# Protective Equipment:

- Lab coat

- Protective wrap-around goggles

# Engineering Controls:

Setup and operation is to be conducted on an open laboratory bench that is well ventilated. A grounded and surge-protected circuit must be utilized to protect the electronics of the rotary evaporator. The machine should be regularly checked and joints re-greased to minimize leaks and the chance of implosion. Ensure that all glassware is free of cracks and star cracks.  Ensure that the bump trap and distilling flask are clamped together using plastic Keck clips.

# Special Handling or Storage Requirements:

Delicate glassware is utilized and needs to be handled with care. The glassware must be identified for any type of crack to minimize the chance of implosion. When the machine is not in use, ensure that solvents are not left in the traps.  If solvent bumps into the bump trap, rinse the bump trap with solvent and return these rinses to the original flask.  If solvent bumps into the condenser, take apart the condenser and vacuum tube apparatus and rinse everything out with acetone.  Discard these acetone rinses into the organic solvent waste container.

# Spill Response Procedure:

Consult the SDS for the chemical spilled. If the chemical spilled is in the water bath, recover the water from the bath, rinse out the bath thoroughly and return the bath to the machine. If the cooling bath leaks, clean with disposable paper towel and place in the garbage.

# Decontamination/Waste Disposal:

The solvent collected by the machine should be disposed of in the proper waste container. The collected solvent should be disposed of immediately after utilizing the machine, and the solvent trap emptied after each use.

# Special Emergency Procedures:

If a major spill or glass implosion should occur, contact the lab PI for proper cleanup instructions.

Procedure may be performed after normal working hours (M-F, 8-5)?	Yes    No

Procedure may be left unattended?		Yes     No

# Step by Step Procedure:

# Preparing the Rotovap for use

- Prior to use, inspect the machine for cracks, or places that look corroded and not structurally sound. It is important to also check the joints to ensure that they are well greased and will seal.  If the machine looks intact, turn on the power to the recirculating chiller and allow the temperature of the antifreeze/water mixture to cool to 4 °C.

- While the machine is cooling, check the solvent traps and empty them into the corresponding waste containers. It is also necessary to clean the bump trap with acetone to ensure that no residue from a prior use remains.

- Next fill the hot water bath if the level has diminished with deionized water, and set the appropriate temperature for it. Most organic solvents will evaporate at 30-40 °C.

- Slowly add dry ice to the isopropanol in the cold finger until the amount of bubbling calms down.

Removing most solvents (boiling point less than 100 °C) using the rotovap

- Step 6 may also be done before step 5, if desired.  Connect the round-bottomed flask containing the solvent to be evaporated to the bump trap and secure with a plastic Keck clip.

- Connect the bump trap to the rotovap and secure this joint using the plastic/metal clip on the rotovap as well as a Keck clip.

- Lower the flask into the water bath with the solvent level equal to the water bath level.  Do not fill the flask more than 2/3 full, especially if using a solvent that is prone to excessive bumping (boiling), such as CH2Cl2, EtOAc, or water.

- Turn on the motor to rotate the flask.  Half of maximum speed is usually fine.

- Before turning on the vacuum pump connected to the rotovap, rotate the black vacuum control knob counterclockwise so it feels loose.

- Ensure that all valves on the rotovap are closed to the atmosphere, and turn on the vacuum pump.

- Slowly increase the vacuum by rotating the black vacuum control knob clockwise until dripping is visible from either the cooling condenser or the cold finger.  If the distilling flask is bubbling vigorously or out of control, immediately reduce the vacuum in one (or more) of four ways:  1) turn the black vacuum control knob counterclockwise, 2) flipping the black “gas ballast” switch on the vacuum pump to release about 10 torr of pressure, 3) rotate the glass valve attached to the cooling condenser until hissing is audible, or 4) loosen the black knob at the top of the cooling condenser until hissing is audible.  The flask may also be raised out of the hot water bath in combination with one of these vacuum-reducing methods. Controlling the rate of solvent evaporation is crucial; it must not be done too quickly.  Visible bubbling may or may not occur, but this is immaterial.

Note on removing solvents with a boiling point over 100 °C

If solvent condenses in the bump trap, increase vacuum or gently heat the bump trap with a heat gun. If a heat gun is used, remove the Keck clip; the vacuum will hold the glassware together.  Also use the lowest temperature setting on the heat gun that is necessary to chase the solvent from the bump trap into the condenser.

- Once all the solvent has evaporated, turn off the vacuum pump, stop rotating the flask, and raise the flask out of the water bath.

- To alleviate removal of the flask while the vacuum completes its mandatory cleaning cycle, slowly open one of the vent valves (items 3 or 4 in step 11 above), then remove the flask (and the bump trap, if necessary).

- Empty both solvent traps into the appropriate waste containers, and clean the bump trap by rinsing it with acetone. Discard these acetone rinses in the organic waste container.

Approved By (PI) : Stephen Chamberland, Ph.D.					Date: 8/29/12

PI Signature: ___________________________________________________		Date: ______________

EH&S Signature: ________________________________________________		Date: ______________